## Front Page
Test Start Time
Test Finish Time
|  | Total Test Duration | 0:00
Customer Details
Test Date: | Certification Number:
Company: | Plant Number
Model:
Serial Number:
Case Number:
Unit 2 Viceroy Court Bedford Road,
Petersfield, Hampshire GU32 3LJ T: 01483 302700 E: info@caltest.co.uk www.caltest.co.uk
Pacific Unit Automated Test

## Checks (hidden)
Test Start Time
Test Finish Time
[5 empty] | Total Test Duration | 0
Customer Details |  |  | Test Equipment
Test Date: |  |  | Manufacture: |  | N4L
Company: |  |  | Model: |  | PPA5500
Model: | 0 |  | Serial Number:
Serial Number: |  |  | Certification Number:
Case Number: |  |  | Calibration Recal Date:
Type | Unit | Test | Controller type | Type
3ph | 0 | 0 | #N/A
3pht | 0 | 0
3ph AFX | 0 | 0
2ph | 0 | 0
2pht | 0 | 0
115ASX | 0 | 0
115ASXT | 0 | 0
Unit 2 Viceroy Court Bedford Road,
Petersfield, Hampshire GU32 3LJ T: 01483 302700 E: info@caltest.co.uk www.caltest.co.uk
Pacific Unit Automated Test

## 3 Ph
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 6 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | K. Gawlik
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | #N/A
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | #N/A
Results are recorded following repair and adjustments.
Calibration reference: |  | #N/A
#N/A
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 6 Pages
3 Phase Voltage Output and Metering.
Test Method:
The 3 phase output and metering accuracy were calibrated by selecting the 3 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded.
#N/A
Frequenc 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
B-N | 10 |  |  |  | #N/A |  |  |  | #N/A
C-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
B-N | 25 |  |  |  | #N/A |  |  |  | #N/A
C-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
B-N | 50 |  |  |  | #N/A |  |  |  | #N/A
C-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
B-N | 75 |  |  |  | #N/A |  |  |  | #N/A
C-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
B-N | 100 |  |  |  | #N/A |  |  |  | #N/A
C-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
B-N | 115 |  |  |  | #N/A |  |  |  | #N/A
C-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
B-N | 125 |  |  |  | #N/A |  |  |  | #N/A
C-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
B-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
C-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 6 Pages
2 Phase Voltage Output and Metering.
Test Method:
The 2 phase output and metering accuracy were calibrated by selecting the 2 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and B.
#N/A
Frequency 50Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 20 |  |  |  | #N/A |  |  |  | #N/A
A-B | 50 |  |  |  | #N/A |  |  |  | #N/A
A-B | 100 |  |  |  | #N/A |  |  |  | #N/A
A-B | 150 |  |  |  | #N/A |  |  |  | #N/A
A-B | 200 |  |  |  | #N/A |  |  |  | #N/A
A-B | 230 |  |  |  | #N/A |  |  |  | #N/A
A-B | 250 |  |  |  | #N/A |  |  |  | #N/A
A-B | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 6 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave
Frequency 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | #N/A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 6 Pages
3 Phase Frequency Response.
Test Method:
The 3 phase frequency response was calibrated by selecting the 3 phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, the measured output voltage and frequency were recorded.
The UUT settings were as follows: Form 3, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 1000 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 6 of 6 Pages
3 Phase Current Metering.
Test Method:
The 3 phase current metering accuracy was calibrated by selecting the 3 phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, the displayed current readings were recorded.
The UUT settings were as follows: Form 3, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 60Hz.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
1 Phase Current Metering.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Ki | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## 3 Ph Tx
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 8 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | I. A. Norman
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Magnetic module s/n |  |  | 5647
Transformer ratio |  |  | 2
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply
Results are recorded following adjustments.
Calibration reference:
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 8 Pages
3 Phase Voltage Output and Metering.
Test Method:
The 3 phase output and metering accuracy were calibrated by selecting the 3 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded.
#N/A
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 | [7 empty] | #N/A
B-N | 10 | [7 empty] | #N/A
C-N | 10 | [7 empty] | #N/A
A-N | 25 | [7 empty] | #N/A
B-N | 25 | [7 empty] | #N/A
C-N | 25 | [7 empty] | #N/A
A-N | 50 | [7 empty] | #N/A
B-N | 50 | [7 empty] | #N/A
C-N | 50 | [7 empty] | #N/A
A-N | 75 | [7 empty] | #N/A
B-N | 75 | [7 empty] | #N/A
C-N | 75 | [7 empty] | #N/A
A-N | 100 | [7 empty] | #N/A
B-N | 100 | [7 empty] | #N/A
C-N | 100 | [7 empty] | #N/A
A-N | 115 | [7 empty] | #N/A
B-N | 115 | [7 empty] | #N/A
C-N | 115 | [7 empty] | #N/A
A-N | 125 | [7 empty] | #N/A
B-N | 125 | [7 empty] | #N/A
C-N | 125 | [7 empty] | #N/A
A-N | #N/A | [7 empty] | #N/A
B-N | #N/A | [7 empty] | #N/A
C-N | #N/A | [7 empty] | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 8 Pages
3 Phase Voltage Output and Metering Transformer Mode.
Test Method:
The 3 phase output and metering accuracy were calibrated by selecting the 3 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded.
The UUT settings were as follows: Form 3, Ext Sense, CSC On, Tx Mode, No Load applied, Sine Wave,
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 20 |  |  |  | #N/A |  |  |  | #N/A
B-N | 20 |  |  |  | #N/A |  |  |  | #N/A
C-N | 20 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
B-N | 50 |  |  |  | #N/A |  |  |  | #N/A
C-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
B-N | 100 |  |  |  | #N/A |  |  |  | #N/A
C-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 150 |  |  |  | #N/A |  |  |  | #N/A
B-N | 150 |  |  |  | #N/A |  |  |  | #N/A
C-N | 150 |  |  |  | #N/A |  |  |  | #N/A
A-N | 200 |  |  |  | #N/A |  |  |  | #N/A
B-N | 200 |  |  |  | #N/A |  |  |  | #N/A
C-N | 200 |  |  |  | #N/A |  |  |  | #N/A
A-N | 230 |  |  |  | #N/A |  |  |  | #N/A
B-N | 230 |  |  |  | #N/A |  |  |  | #N/A
C-N | 230 |  |  |  | #N/A |  |  |  | #N/A
A-N | 250 |  |  |  | #N/A |  |  |  | #N/A
B-N | 250 |  |  |  | #N/A |  |  |  | #N/A
C-N | 250 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
B-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
C-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 8 Pages
2 Phase Voltage Output and Metering.
Test Method:
The 2 phase output and metering accuracy were calibrated by selecting the 2 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and B.
#N/A
Frequency 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 20 |  |  |  | #N/A |  |  |  | #N/A
A-B | 50 |  |  |  | #N/A |  |  |  | #N/A
A-B | 100 |  |  |  | #N/A |  |  |  | #N/A
A-B | 150 |  |  |  | #N/A |  |  |  | #N/A
A-B | 200 |  |  |  | #N/A |  |  |  | #N/A
A-B | 230 |  |  |  | #N/A |  |  |  | #N/A
A-B | 250 |  |  |  | #N/A |  |  |  | #N/A
A-B | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 8 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | #N/A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 6 of 8 Pages
1 Phase Voltage Output and Metering Transformer Mode.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Ext Sense, CSC On, Tx Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 20 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 150 |  |  |  | #N/A |  |  |  | #N/A
A-N | 200 |  |  |  | #N/A |  |  |  | #N/A
A-N | 230 |  |  |  | #N/A |  |  |  | #N/A
A-N | 250 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 7 of 8 Pages
3 Phase Frequency Response.
Test Method:
The 3 phase frequency response was calibrated by selecting the 3 phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, the measured output voltage and frequency were recorded.
The UUT settings were as follows: Form 3, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | 1000 |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
B-N | [8 empty] | #N/A
C-N | [8 empty] | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 8 of 8 Pages
3 Phase Current Metering.
Test Method:
The 3 phase current metering accuracy was calibrated by selecting the 3 phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, the displayed current readings were recorded.
The UUT settings were as follows: Form 3, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 400Hz.
Phase |  | Measured Current |  |  | Displayed Current |  | Displayed Current Tx |  | Displayed Current Limits
|  | (A) |  |  | (A) |  | (A) |  | (± A)
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
A-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
B-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
C-N |  | #N/A |  |  |  |  | #N/A |  | #N/A
1 Phase Current Metering.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N (Direct) |  | #N/A | [5 empty] | #N/A
A-N (Tx) |  | #N/A | [5 empty] | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Kin | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## 3 Ph AFX
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 11 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | K. Gawlik
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source | 1243
Serial Number
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | A002-2, A003-1, A004-2, A008, C007
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | 400V -6%/+10%, 50 Hz ±1%
Results are recorded following repair and adjustments.
Calibration reference: |  | AFX Series operational manual, P/N 160620-10, Rev 1.1.7. 2018
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 11 Pages
3 Phase AC Output Voltage Linearity and Metering.
Test Method:
The 3 phase output and metering accuracy was calibrated by selecting the 3 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All phases were measured with respect to Neutral. The UUT settings were as follows: Mode AC, Form 3, Int sense, CSC on, High Range, No Load applied. Frequency 400 Hz Sine Wave * Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 10 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 10 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 25 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 25 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 25 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 50 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 50 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 50 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 75 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 75 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 75 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 100 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 100 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 100 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 115 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 115 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 115 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 125 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 125 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 125 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 150 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 150 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 150 |  |  |  | 0.875 |  |  |  | 0.75
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 11 Pages
3 Phase AC Output Voltage Linearity and Metering Cont:
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 200 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 200 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 200 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 240 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 240 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 240 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 270 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 270 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 270 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 300 |  |  |  | 0.875 |  |  |  | 0.75
B-N | 300 |  |  |  | 0.875 |  |  |  | 0.75
C-N | 300 |  |  |  | 0.875 |  |  |  | 0.75
Uncertainties |  |  | AC Voltage ± 10 mV
2 Phase AC Output Voltage Linearity and Metering Cont:
Test Method:
The 2 phase output and metering direct coupled accuracy was calibrated by selecting the 2 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and Phase B. The UUT settings were as follows: Mode AC, Form 2, Int sense, CSC on, High Range, No Load applied. Frequency 400 Hz Sine Wave. * Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 50 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 100 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 150 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 200 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 300 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 400 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 500 |  |  |  | 1.5 |  |  |  | 1.5
A-B | 600 |  |  |  | 1.5 |  |  |  | 1.5
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 11 Pages
1 Phase AC Output Voltage Linearity and Metering.
Test Method:
The 1 phase output and metering direct coupled accuracy was calibrated by selecting the 1 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. For single phase operation, all three phase outputs were connected in parallel configuration. The UUT settings were as follows: Mode AC, Form 1, Int sense, CSC on, High Range, No Load applied. Frequency 400 Hz Sine Wave. * Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 25 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 50 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 75 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 100 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 115 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 125 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 150 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 200 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 240 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 270 |  |  |  | 0.875 |  |  |  | 0.75
A-N | 300 |  |  |  | 0.875 |  |  |  | 0.75
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 11 Pages
3 Phase Frequency Response.
Test Method:
The 3 phase voltage frequency response was calibrated by selecting the 3 Phase operation mode, the output voltages were selected for 115 V and the frequency was set to the values below, the measured Output Voltage and Frequency were recorded. All phases were measured with respect to Neutral. The UUT settings were as follows: Mode AC, Form 3, Int sense, CSC on, High Range, No Load applied. Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | 0.005 |  |  |  | 0.75
B-N | [8 empty] | 0.75
C-N | [8 empty] | 0.75
A-N | 60 |  |  |  | 0.006 |  |  |  | 0.75
B-N | [8 empty] | 0.75
C-N | [8 empty] | 0.75
A-N | 400 |  |  |  | 0.04 |  |  |  | 0.75
B-N | [8 empty] | 0.75
C-N | [8 empty] | 0.75
A-N | 1000 |  |  |  | 0.1 |  |  |  | 0.75
B-N | [8 empty] | 0.75
C-N | [8 empty] | 0.75
A-N | 1200 |  |  |  | 0.12 |  |  |  | 0.75
B-N | [8 empty] | 0.75
C-N | [8 empty] | 0.75
Uncertainties |  |  | AC Voltage ± 50 mV
Frequency 50 Hz to 1.2 kHz ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 6 of 11 Pages
AC 3 Phase Current Metering.
Test Method:
The 3 phase Current metering accuracy was calibrated by selecting the 3 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, the displayed current readings were recorded. All phases were measured with respect to Neutral. The UUT settings were as follows: Mode AC, Form 3, Int sense, CSC off, High Range, Frequency 400 Hz Sine Wave.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
1 Phase Current Metering.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
Uncertainties |  | Measured AC Current Up to 20 A ± 20 mA
20 A to 120 A ± 60 mA
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 7 of 11 Pages
AC 3 Phase Load Regulation
Test Method:
The Load Regulation was calibrated by setting the Unit Under Test to supply a nominal 115 V output with a nominal current of 0 A flowing into the load, the output voltage value was recorded.
The load was adjusted for a maximum current output and the output voltage was again recorded. The UUT settings were as follows: Mode AC, Form 3, Ext sense, CSC on, High Range, Sine Wave, 400Hz.
Phase |  | Voltage |  | Load current |  | Measured Voltage |  | Voltage Limits
|  | (V) |  | (A) |  | (V) |  | (± V)
A-N |  | 115 |  | 0 |  |  |  | Ref
B-N |  | 115 |  | 0 |  |  |  | Ref
C-N |  | 115 |  | 0 |  |  |  | Ref
A-N |  | 115 |  | #N/A |  |  |  | 0.023 V from Ref
B-N |  | 115 |  | #N/A |  |  |  | 0.023 V from Ref
C-N |  | 115 |  | #N/A |  |  |  | 0.023 V from Ref
Uncertainties |  |  | AC Voltage ± 50 mV
DC OUTPUT
DC 3 Phase Load Regulation
Test Method:
The Load Regulation was calibrated by setting the Unit Under Test to supply a nominal 170 V output with a nominal current of 0 A flowing into the load, the output voltage value was recorded.
The load was adjusted for a maximum current output and the output voltage was again recorded. The UUT settings were as follows: Mode DC, Form 3, Ext sense, CSC on, High Range.
Phase |  | Voltage |  | Load current |  | Measured Voltage |  | Voltage Limits
|  | (V) |  | (A) |  | (V) |  | (± V)
A-N |  | 170 |  | 0 |  |  |  | Ref
B-N |  | 170 |  | 0 |  |  |  | Ref
C-N |  | 170 |  | 0 |  |  |  | Ref
A-N |  | 170 |  | #N/A |  |  |  | 0.034 V from Ref
B-N |  | 170 |  | #N/A |  |  |  | 0.034 V from Ref
C-N |  | 170 |  | #N/A |  |  |  | 0.034 V from Ref
Uncertainties |  |  | DC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 8 of 11 Pages
3 Phase DC Output Voltage Linearity and Metering.
Test Method:
The 3 phase output and metering accuracy was calibrated by selecting the 3 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All phases were measured with respect to Neutral. The UUT settings were as follows: Mode DC, Form 3, Int sense, CSC on, High Range, No Load applied. * Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 25 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 25 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 25 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 50 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 50 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 50 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 75 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 75 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 75 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 100 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 100 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 100 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 120 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 120 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 120 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 150 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 150 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 150 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 200 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 200 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 200 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 250 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 250 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 250 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
Uncertainties |  |  | DC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 9 of 11 Pages
3 Phase DC Output Voltage Linearity and Metering Cont:
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 300 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 300 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 300 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 350 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 350 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 350 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 400 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 400 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 400 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 425 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
B-N | 425 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
C-N | 425 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
Uncertainties |  |  | DC Voltage ± 10 mV
2 Phase DC Output Voltage Linearity and Metering Cont:
Test Method:
The 2 phase output and metering direct coupled accuracy was calibrated by selecting the 2 Phase direct coupled operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and Phase B. The UUT settings were as follows: Mode DC, Form 2, Int sense, CSC on, High Range , No Load applied.
* Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 100 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 200 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 300 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 400 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 500 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 600 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 700 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 800 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
A-B | 850 |  | #REF! |  | 2.2 |  | #REF! |  | 2.125
Uncertainties |  |  | DC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 10 of 11 Pages
1 Phase DC Output Voltage Linearity and Metering.
Test Method:
The 1 phase output and metering accuracy was calibrated by selecting the 1 Phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. For single phase operation, all three phase outputs were connected in parallel configuration. The UUT settings were as follows: Mode DC, Form 1, Int sense, CSC on, High Range, No Load applied. * Display Limits refer to Limits between Displayed output and Measured value.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 25 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 50 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 100 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 150 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 200 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 250 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 300 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 350 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 400 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
A-N | 425 |  | #REF! |  | 1.1 |  | #REF! |  | 1.0625
Uncertainties |  |  | DC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 11 of 11 Pages
3 DC Phase Current Metering.
Test Method:
The 3 phase Current metering accuracy was calibrated by selecting the 3 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, the displayed current readings were recorded. All phases were measured with respect to Neutral. The UUT settings were as follows: Mode DC, Form 3, Int sense, CSC off, High Range.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
A-N |  | #N/A | [5 empty] | #N/A
B-N |  | #N/A | [5 empty] | #N/A
C-N |  | #N/A | [5 empty] | #N/A
1 Phase Current Metering.
Phase |  | Measured Current |  |  | Displayed Current |  |  | Displayed Current Limits
|  | (A) |  |  | (A) |  |  | (± A)
A-N |  | #N/A | [5 empty] | #N/A
Uncertainties |  | Measured DC Current Up to 20 A ± 20 mA
20 A to 100 A ± 30 mA
END

## 2 Ph
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 5 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | I. A. Norman
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | #N/A
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | #N/A
Results are recorded following adjustments.
Calibration reference: |  | #N/A
#N/A
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 5 Pages
2 Phase Voltage Output and Metering.
Test Method:
The 2 phase output and metering accuracy were calibrated by selecting the 2 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and B.
#N/A
Frequency 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 20 |  |  |  | #N/A |  |  |  | #N/A
A-B | 50 |  |  |  | #N/A |  |  |  | #N/A
A-B | 100 |  |  |  | #N/A |  |  |  | #N/A
A-B | 150 |  |  |  | #N/A |  |  |  | #N/A
A-B | 200 |  |  |  | #N/A |  |  |  | #N/A
A-B | 230 |  |  |  | #N/A |  |  |  | #N/A
A-B | 250 |  |  |  | #N/A |  |  |  | #N/A
A-B | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 5 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
#N/A
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | #N/A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 5 Pages
Frequency Response.
Test Method:
The Single phase Frequency response was calibrated by selecting the 1 Phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, The measured Output Voltage and Frequency were recorded.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 5 Pages
Current Metering.
Test Method:
The Current metering Accuracy was calibrated by selecting the 1 Phase operation then the 2 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, The displayed current readings were recorded.
The UUT settings were as follows: Form 1 or 2, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 400Hz.
Measured Current |  | Displayed Current 1 Phase |  | Measured Current |  | Displayed Current 2 Phase |  | Displayed Current Limits
(A) |  | (A) |  | (A) |  | (A) |  | (± A)
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
#N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Kin | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## 2 Ph Tx
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 6 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | I. A. Norman
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Magnetic module s/n
Transformer ratio |  |  | 2
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | #N/A
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | #N/A
Results are recorded following adjustments.
Calibration reference: |  | #N/A
#N/A
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 6 Pages
2 Phase Voltage Output and Metering.
Test Method:
The 2 phase output and metering accuracy were calibrated by selecting the 2 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and B.
The UUT settings were as follows: Form 2, Int Sense, CSC off, Direct Mode, No Load applied, Sine Wave,
Frequency 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 20 |  |  |  | #N/A |  |  |  | #N/A
A-B | 50 |  |  |  | #N/A |  |  |  | #N/A
A-B | 100 |  |  |  | #N/A |  |  |  | #N/A
A-B | 150 |  |  |  | #N/A |  |  |  | #N/A
A-B | 200 |  |  |  | #N/A |  |  |  | #N/A
A-B | 230 |  |  |  | #N/A |  |  |  | #N/A
A-B | 250 |  |  |  | #N/A |  |  |  | #N/A
A-B | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 6 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | #N/A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 6 Pages
1 Phase Voltage Output and Metering Transformer Mode.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Ext Sense, CSC On, Tx Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 20 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 150 |  |  |  | #N/A |  |  |  | #N/A
A-N | 200 |  |  |  | #N/A |  |  |  | #N/A
A-N | 230 |  |  |  | #N/A |  |  |  | #N/A
A-N | 250 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 6 Pages
Frequency Response.
Test Method:
The Single phase Frequency response was calibrated by selecting the 1 Phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, The measured Output Voltage and Frequency were recorded.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
A-N | 1000 |  |  |  | #N/A |  |  |  | #N/A
A-N | #N/A |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 6 of 6 Pages
Current Metering.
Test Method:
The Current metering Accuracy was calibrated by selecting the 1 Phase operation then the 2 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, The displayed current readings were recorded.
The UUT settings were as follows: Form 1 or 2, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 400Hz.
Measured Current | Displayed Current 1 Phase |  | Measured Current | Displayed Current 2 Phase |  | Measured Current | Displayed Current 1 Phase Tx |  | Displayed Current Limits
(A) | (A) |  | (A) | (A) |  | (A) | (A) |  | (± A)
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
#N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Kin | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## 115ASX
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | C024567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 4 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | I. A. Norman
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | A002-2, A008, C007
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | 230V -6%/+10%, 50 Hz ±1%
Results are recorded following adjustments.
Calibration reference: |  | #N/A
#N/A
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | C024567
Page 2 of 4 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Int Sense, CSC off, Direct Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | 132 |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | @ 5A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | C024567
Page 3 of 4 Pages
Frequency Response.
Test Method:
The Single phase Frequency response was calibrated by selecting the 1 Phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, The measured Output Voltage and Frequency were recorded.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
A-N | 1000 |  |  |  | #N/A |  |  |  | #N/A
A-N | 1200 |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | C024567
Page 4 of 4 Pages
Current Metering.
Test Method:
The Current metering Accuracy was calibrated by selecting the 1 Phase operation then the 2 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, The displayed current readings were recorded.
The UUT settings were as follows: Form 1, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 400Hz.
Measured Current |  |  |  | Displayed Current 1 Phase |  |  |  | Displayed Current Limits
(A) |  |  |  | (A) |  |  |  | (± A)
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
#N/A | [7 empty] | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Kin | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## 115ASXT
CERTIFICATE OF CALIBRATION
ISSUED BY: |  | Caltest Instruments Ltd
DATE OF ISSUE: |  | 2020-12-10 |  |  |  | Certificate Number |  |  | 04567
[5 empty] | Unit 2 Viceroy Court |  |  |  | Page 1 of 5 Pages
Bedford Road
[5 empty] | Petersfield |  |  |  | Approved Signatory
Hampshire
GU32 3LJ
[5 empty] | Tel 01483 302700 |  |  |  | I. A. Norman
Customer: |  |  | Caltest Instruments Ltd
Unit 2 Viceroy Court
Bedford Road
Petersfield
Hampshire
GU32 3LJ
Customer Order No. |  |  | 78546978
Case No. |  |  | C024567 | A
Manufacturer |  |  | Pacific
Model Number
Equipment Description |  |  | AC Power Source
Serial Number |  |  | 1243
Transformer ratio |  |  | 2
Ident/Plant Number |  |  | CT125
Date of Calibration |  |  | 2020-12-10
Procedures Used |  |  | A002-2, A008, C007
Ambient Temperature |  |  | 23 °C ± 1 °C
Relative Humidity |  |  | 50 % rh ± 15 % rh
Nominal Mains Supply |  |  | 230V -6%/+10%, 50 Hz ±1%
Results are recorded following adjustments.
Calibration reference: |  | #N/A
#N/A
Measurements were made following a minimum stabilisation period of 2 hours in the Laboratory environment.
All measurements were made using equipment traceable to National or International Standards.
The reported expanded uncertainty is based on a standard uncertainty multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. The quoted uncertainty refers only to the measured value and does not carry any implication regarding the long term stability of the instrument.
Standards used:
Calibrated By: |  | I. A. Norman
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 2 of 5 Pages
1 Phase Voltage Output and Metering.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Int Sense, CSC Off, Direct Mode, No Load applied, Sine Wave
Frequency 400Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-N | 10 |  |  |  | #N/A |  |  |  | #N/A
A-N | 25 |  |  |  | #N/A |  |  |  | #N/A
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 75 |  |  |  | #N/A |  |  |  | #N/A
A-N | 100 |  |  |  | #N/A |  |  |  | #N/A
A-N | 115 |  |  |  | #N/A |  |  |  | #N/A
A-N | 125 |  |  |  | #N/A |  |  |  | #N/A
A-N | 132 |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Other Checks
Current Limit Checked |  |  |  |  | ü | @ 5A
External Sense Function Checked |  |  |  |  | ü
Power Meter Function Checked |  |  |  |  | ü
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 3 of 5 Pages
Transformer Coupled Output.
Test Method:
The 1 phase output and metering accuracy were calibrated by selecting the 1 phase operation mode, the following output voltages were selected, the measured output voltage and the value of the output meter were recorded. All measurements were made between Phase A and N.
The UUT settings were as follows: Form 1, Ext Sense, CSC on, Tx Mode, No Load applied, Sine Wave,
Frequency 60Hz.
Phase | Setting |  | Display |  | Display Limits |  | Measured Value |  | Measured Value Limits
| (V) |  | (V) |  | (± V) |  | (V) |  | (± V)
A-B | 20 |  |  |  | #N/A |  |  |  | #N/A
A-B | 50 |  |  |  | #N/A |  |  |  | #N/A
A-B | 100 |  |  |  | #N/A |  |  |  | #N/A
A-B | 150 |  |  |  | #N/A |  |  |  | #N/A
A-B | 200 |  |  |  | #N/A |  |  |  | #N/A
A-B | 230 |  |  |  | #N/A |  |  |  | #N/A
A-B | 250 |  |  |  | #N/A |  |  |  | #N/A
A-B | 264 |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 4 of 5 Pages
Frequency Response.
Test Method:
The Single phase Frequency response was calibrated by selecting the 1 Phase operation mode, the output voltages were selected for 120V and the frequency was set to the values below, The measured Output Voltage and Frequency were recorded.
The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave.
Phase | Setting |  | Measured Frequency |  | Frequency Limits |  | Measured Voltage |  | Measured Value Limits
| (Hz) |  | (Hz) |  | (± Hz) |  | (V) |  | (± V)
A-N | 50 |  |  |  | #N/A |  |  |  | #N/A
A-N | 60 |  |  |  | #N/A |  |  |  | #N/A
A-N | 400 |  |  |  | #N/A |  |  |  | #N/A
A-N | 1000 |  |  |  | #N/A |  |  |  | #N/A
A-N | 1200 |  |  |  | #N/A |  |  |  | #N/A
Uncertainties |  |  | AC Voltage ± 10 mV
Frequency ± 0.001 Hz
CERTIFICATE OF CALIBRATION | [7 empty] | Certificate Number | 04567
Page 5 of 5 Pages
Current Metering.
Test Method:
The Current metering Accuracy was calibrated by selecting the 1 Phase operation then the 2 Phase operation mode, a resistive load was connected to each phase in turn, the load was adjusted for the following currents, The displayed current readings were recorded.
The UUT settings were as follows: Form 1 or 2, Int sense, CSC off, Direct Mode, Sine Wave,
Frequency 400Hz.
Measured Current |  | Displayed Current 1 Phase |  | Measured Current Tx mode |  | Displayed Current Tx mode |  | Displayed Current Limits
(A) |  | (A) |  | (A) |  | (A) |  | (± A)
0.5 |  |  |  | 0.5 |  |  |  | #N/A
1 |  |  |  | 1 |  |  |  | #N/A
2 |  |  |  | 2 |  |  |  | #N/A
4 |  |  |  | 4 |  |  |  | #N/A
6 |  |  |  | 6 |  |  |  | #N/A
8 |  |  |  | N/A |  |  |  | #N/A
10 |  |  |  | N/A |  |  |  | #N/A
15 |  |  |  | N/A |  |  |  | #N/A
Uncertainties |  | Displayed AC Current ± 1 Display Digit
Measured AC Current ± 1 mA
END
Pacific AC Power Source K Factors
Customer | Caltest Instruments Ltd |  |  |  | Ser No | 1243 |  | Controller: |  | #N/A
Model |  |  |  |  | Date | 2020-12-10 |  | Type: |  | #N/A
(Fn 4, 3, ) K Factor display UPC, Form 3
Function | Phase | K Factor Display |  | Phase | K Factor Display |  | Phase | K Factor Display
Kin | A: |  |  | B: |  |  | C:
Kex | A: |  |  | B: |  |  | C:
Kin | A: |  |  | B: |  |  | C:
Ko | A: |  |  | B: |  |  | C:
| (Fn 4, 3, ) K Factor display UPC3, Form 2 |  |  |  |  | (Fn 4, 3, ) K Factor display UPC3, Form 1
| Function |  |  |  |  | Function
| Kin |  |  |  |  | Kin
| Kex |  |  |  |  | Kex
| Kin |  |  |  |  | Kin
| Ko |  |  |  |  | Ko
New K Factor Display = Old K Factor * (Test Equipment Value/UPC Meter Display)
(Fn 3,3)
GPIB PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
GPIB ADDRESS
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
SERIAL PORT
UPC STATUS: |  |  |  |  | FIRMWARE: |  |  |  | V
BAUD: |  |  |  |  | PARITY:
TRANSFORMER RATIO |  |  |  |  | PAs=
AMPS: VOLTS RATIO |  |  |  |  | A=1 V |  |  |  | SEC
Note on E25: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For Avionics choose A

## Test Equ (hidden)
Test Equipment
Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033.
Above 20A | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020.
Above 20A | Cropico RS2 Standard Resistor 0.1Ω Ser No.19420. Calibrated 08th June 2020.
Newtons 4th PPA5530 Power Analyser, Serial number 165-00791. Calibrated 28th January. Certificate number 165-00791.
Above 50A | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020.
Newtons 4th PPA530 Power Analyser, Serial number 115-01543. Calibrated 31st January 2020. UKAS Certificate number 115-01543310120.
Newton’s 4th HF500 Current Shunt Serial number 1979. Calibration certificate number 04626, Calibrated 16th January 2020.
Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020.
Manuals:
AMX | AMX operation manual, part number 139250, Revision B. Dated April 2012.
ASX | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017.
ACX | ACX-Series operational manual, part number 150118-10, sixth edition, June 2014.
AFX | AFX Series operational manual, P/N 160620-10, Rev 1.1.7. 2018
ADX | ADX-Series operational manual, part number 150180-10, sixth edition, February 2016.
LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019.
UPC1 | UPC-1/3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015.
UPC12 | UPC-12/32 operation manual, part number 133620, Sixth Edition, dated April 2020.
UPC3 | UPC-1/3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015.
UPC31
UPC32 | UPC-12/32 operation manual, part number 133620, Sixth Edition, dated April 2020.
Procedures:
A004-2 | Current 20A-100A
A008 | Frequency
A002-2 | Using Fluke
C007
Under 20A | A002-2, A008, C007
Above 20A | A002-2, A004-2, A008, C007
DCA up 100A | A002-2, A003-1, A004-2, A008, C007
Users:
I. A. Norman
K. Gawlik
D. Holyoake
Input
230V -6%/+10%, 50 Hz ±1%
400V -6%/+10%, 50 Hz ±1%
Onsite condition
Results:
Results are recorded “As Received” values.
Results are recorded following repair and adjustments.
Results are recorded following adjustments.
3ph V | The UUT settings were as follows: Form 3, Int Sense, CSC off, Direct Mode, No Load applied, Sine Wave,
LMX | The UUT settings were as follows: Form 3, Int Sense, CSC on, Direct Mode, No Load applied, Sine Wave,
2 ph V | The UUT settings were as follows: Form 2, Int Sense, CSC off, Direct Mode, No Load applied, Sine Wave,
LmX | The UUT settings were as follows: Form 2, Int Sense, CSC on, Direct Mode, No Load applied, Sine Wave,
1p v | The UUT settings were as follows: Form 1, Int Sense, CSC Off, Direct Mode, No Load applied, Sine Wave
Lmx | The UUT settings were as follows: Form 1, Int Sense, CSC On, Direct Mode, No Load applied, Sine Wave

## 3Ph data (hidden)
Model | Unit Manual | UPC Manual | Input | Procedure | Equipment | EQU2 | EQU3 | EQU4 |  | Controller | Type | Model | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current Single | MAX 3 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
305AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | AMX | 305AMX-UPC3 | 135 | 1200 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
305AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | AMX | 305AMX-UPC32 | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
305LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | LMX | LMX | 305LMX | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
308AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | AMX | 308AMX-UPC3 | 135 | 1200 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
308AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | AMX | 308AMX-UPC32 | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
308LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | LMX | LMX | 308LMX | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
312AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | AMX | 312AMX-UPC3 | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
312AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | AMX | 312AMX-UPC32 | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
312LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | LMX | LMX | 312LMX | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
315ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | ASX | 315ASX-UPC3 | 132 | 1200 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 2 |  | 12 | 4 | 12 | 6 | 20 | 60
315ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | ASX | 315ASX-UPC32 | 132 | 1200 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 2 |  | 12 | 4 | 12 | 6 | 20 | 60
320AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | AMX | 320AMX-UPC3 | 135 | 1200 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | AMX | 320AMX-UPC32 | 135 | 5000 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 400V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | LMX | LMX | 320LMX | 135 | 5000 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC3 | ASX | 320ASX-UPC3 | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 7 | 5 |  | 20 | 7 | 20 | 6 | 20 | 60
320ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 230V -6%/+10%, 50 Hz ±1% | A002-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. |  |  | 1 | UPC32 | ASX | 320ASX-UPC32 | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 7 | 5 |  | 20 | 7 | 20 | 6 | 20 | 60
345AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC3 | AMX | 345AMX-UPC3 | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC32 | AMX | 345AMX-UPC32 | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | LMX | LMX | 345LMX | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC3 | ASX | 345ASX-UPC3 | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 16 | 36 | 12 | 40 | 120
345ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC32 | ASX | 345ASX-UPC32 | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 16 | 36 | 12 | 40 | 120
360AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC3 | AMX | 360AMX-UPC3 | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC32 | AMX | 360AMX-UPC32 | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | LMX | LMX | 360LMX | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC3 | ASX | 360ASX-UPC3 | 132 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 12 | 40 | 120
360ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. |  | 1 | UPC32 | ASX | 360ASX-UPC32 | 132 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 12 | 40 | 120
390AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC3 | AMX | 390AMX-UPC3 | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC32 | AMX | 390AMX-UPC32 | 135 | 5000 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | LMX | LMX | 390LMX | 135 | 5000 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC3 | ASX | 390ASX-UPC3 | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 60 | 200 | 600
390ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC32 | ASX | 390ASX-UPC32 | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 60 | 200 | 600
3120AMX-UPC3 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC3 | AMX | 3120AMX-UPC3 | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120AMX-UPC32 | AMX operation manual, part number 139250, Revision B. Dated April 2012. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC32 | AMX | 3120AMX-UPC32 | 135 | 5000 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120LMX | LMX Series operation manual, PPS part number 143750-10, Rev 1.0.2, dated 2019. |  | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | LMX | LMX | 3120LMX | 135 | 5000 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120ASX-UPC3 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC-3 operation manual, PPS part number 141020, Fifth Edition, dated March 2015. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC3 | ASX | 3120ASX-UPC3 | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 60 | 200 | 600
3120ASX-UPC32 | ASX Series operation manual, PPS part number 140050, 9th Edition, dated June 2017. | UPC32 operation manual, part number 133620, Sixth Edition, dated April 2020. | 400V -6%/+10%, 50 Hz ±1% | A002-2, A004-2, A008, C007 | Frequency via Quartzlock E8-Y GPS Frequency Reference, Serial number 17012033. | Fluke 8508A Digital Multimeter Serial No. 115260406 UKAS Calibration Certificate No. 0183-075669 dated 06th April 2020. | Cropico RS2 Standard Resistor 0.01Ω Ser No. 6577 Calibrated 08th June 2020. | Newton’s 4th HF100 Current Shunt Serial number 911-11514. Calibrated 09th June 2020. | 1 | UPC32 | ASX | 3120ASX-UPC32 | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 60 | 200 | 600
A004-2 | Current 20A-100A
A008 | Frequency
A002-2 | Using Fluke | [10 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29

## 3PhT data (hidden)
Model | Controller | Type | Procedure | Input | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current Single | MAX 3 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
305AMXT-UPC3 | UPC3 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
305AMXT-UPC32 | UPC32 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
305LMXT | LMX | LMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.3 | 1 |  | 4 | 1.33 | 4 | 6 | 20 | 60
308AMXT-UPC3 | UPC3 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
308AMXT-UPC32 | UPC32 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
308LMXT | LMX | LMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 1 |  | 6 | 2 | 6 | 6 | 20 | 60
312AMXT-UPC3 | UPC3 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
312AMXT-UPC32 | UPC32 | AMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
312LMXT | LMX | LMX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 1 |  | 10 | 3 | 10 | 6 | 20 | 60
315ASXT-UPC3 | UPC3 | ASX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 2 |  | 12 | 4 | 12 | 6 | 20 | 60
315ASXT-UPC32 | UPC32 | ASX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 2 |  | 12 | 4 | 12 | 6 | 20 | 60
320AMXT-UPC3 | UPC3 | AMX | A002-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320AMXT-UPC32 | UPC32 | AMX | A002-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320LMXT | LMX | LMX | A002-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 3 | 4 | 6 | 5 |  | 18 | 6 | 18 | 15 | 50 | 150
320ASXT-UPC3 | UPC3 | ASX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 7 | 5 |  | 20 | 7 | 20 | 6 | 20 | 60
320ASXT-UPC32 | UPC32 | ASX | A002-2, A008, C007 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 7 | 5 |  | 20 | 7 | 20 | 6 | 20 | 60
345AMXT-UPC3 | UPC3 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345AMXT-UPC32 | UPC32 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345LMXT | LMX | LMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 12 | 36 | 15 | 50 | 150
345ASXT-UPC3 | UPC3 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 16 | 36 | 12 | 40 | 120
345ASXT-UPC32 | UPC32 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 12 | 5 |  | 30 | 16 | 36 | 12 | 40 | 120
360AMXT-UPC3 | UPC3 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360AMXT-UPC32 | UPC32 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360LMXT | LMX | LMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 15 | 50 | 150
360ASXT-UPC3 | UPC3 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 12 | 40 | 120
360ASXT-UPC32 | UPC32 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 | 5 |  | 30 | 16 | 48 | 12 | 40 | 120
390AMXT-UPC3 | UPC3 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390AMXT-UPC32 | UPC32 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390LMXT | LMX | LMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 30 | 100 | 300
390ASXT-UPC3 | UPC3 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 60 | 200 | 600
390ASXT-UPC32 | UPC32 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 15 | 24 | 5 |  | 70 | 24 | 72 | 60 | 200 | 600
3120AMXT-UPC3 | UPC3 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120AMXT-UPC32 | UPC32 | AMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120LMXT | LMX | LMX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 30 | 100 | 300
3120ASXT-UPC3 | UPC3 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 60 | 200 | 600
3120ASXT-UPC32 | UPC32 | ASX | A002-2, A004-2, A008, C007 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 10 | 20 | 30 | 5 |  | 90 | 32 | 96 | 60 | 200 | 600
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21

## AFX (hidden)
|  |  |  | AC | [9 empty] | AC | AC | DC |  | DC | DC
|  |  |  | V | VDC | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | spare | MAX 3 PAC | MAX 1 PAC | MAX 3 PDC | MAX 1 PDC | MAX 1 PDC | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6
360AFX | AFX |  |  | 300 | 425 | 1200 | 0.5 | 1 | 2 | 5 | 10 | 15 |  | 16.7 | 50 | 8.3 | 25 | 25 | 0.5 | 1 | 2 | 4 | 6 | 8
390AFX | AFX |  |  | 300 | 425 | 1200 | 1 | 2 | 5 | 10 | 15 | 25 |  | 25 | 75 | 12.9 | 37.9 | 35 | 0.5 | 1 | 2 | 5 | 10 | 12
3120AFX | AFX |  |  | 300 | 425 | 1200 | 1 | 2 | 4 | 10 | 20 | 30 |  | 33.3 | 100 | 16.7 | 50 | 50 | 0.5 | 1 | 2 | 5 | 10 | 15
3150AFX | AFX |  |  | 300 | 425 | 1200 | 1 | 5 | 10 | 20 | 30 | 40 |  | 41.7 | 125 | 21 | 63 | 60 | 1 | 2 | 5 | 10 | 15 | 20
3180AFX | AFX |  |  | 300 | 425 | 1200 | 1 | 5 | 10 | 20 | 35 | 50 |  | 50 | 150 | 27.8 | 75 | 75 | 1 | 2 | 5 | 10 | 15 | 25
3240AFX | AFX |  |  | 300 | 425 | 1200 | 2 | 4 | 8 | 20 | 40 | 65 |  | 66.7 | 200 | 33.4 | 100 | 100 | 1 | 2 | 4 | 10 | 20 | 30
3300AFX | AFX |  |  | 300 | 425 | 1200 | 2 | 10 | 20 | 40 | 60 | 80 |  | 83.3 | 250 | 41.7 | 125 | 125 | 2 | 4 | 10 | 20 | 30 | 40
3450AFX | AFX |  |  | 300 | 425 | 1200 | 3 | 15 | 30 | 60 | 90 | 125 |  | 125 | 375 | 62.5 | 189 | 185 | 3 | 6 | 15 | 30 | 45 | 60
3600AFX | AFX |  |  | 300 | 425 | 1200 | 5 | 10 | 20 | 50 | 100 | 165 |  | 166.7 | 500 | 83.3 | 250 | 250 | 2 | 10 | 20 | 40 | 60 | 80
3750AFX | AFX |  |  | 300 | 425 | 1200 | 5 | 25 | 50 | 100 | 150 | 205 |  | 208.3 | 625 | 104 | 312.5 | 310 | 5 | 10 | 25 | 50 | 75 | 100
3900AFX | AFX |  |  | 300 | 425 | 1200 | 6 | 30 | 60 | 120 | 180 | 250 |  | 250 | 750 | 125 | 375 | 375 | 6 | 12 | 30 | 60 | 90 | 125
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25
|  |  |  | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 13 | 14 | 15 |  | 16 | 17
|  |  |  | 14 |  | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 |  | 28 | 29

## 2Ph data (hidden)
Model | Type | Contr | Input | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current limit | MAX 2 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
105AMX-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
105AMX-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
105LMX | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
108AMX-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
108AMX-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
108LMX | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
110ACX-UPC1 | ACX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5.5 | 10 | 5 | 50
110ADX-UPC1 | ADX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 600 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5.5 | 10 | 5 | 50 |  | ADX
112AMX-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
112AMX-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
112LMX | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
115ACX-UPC1 | ACX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 | 5 | 8 | 15 | 5 | 50
115ADX-UPC1 | ADX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 600 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 | 5 | 8 | 15 | 5 | 50 |  | ADX
118ACX-UPC1 | ACX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 5 | 50
120ASX-UPC1 | ASX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 14 | 20 | 6 | 60
120ASX-UPC12 | ASX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 14 | 20 | 6 | 60
125AMX-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
125AMX-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
125LMX | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
140AMX-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140AMX-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140LMX | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140ASX-UPC1 | ASX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 12 | 120
140ASX-UPC12 | ASX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 12 | 120
160AMX-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160AMX-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160LMX | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160ASX-UPC1 | ASX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 12 | 120
160ASX-UPC12 | ASX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 12 | 120
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19

## 2PhT data (hidden)
Model | Type | Contr | Input | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current limit | MAX 2 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
105AMXT-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
105AMXT-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
105LMXT | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 2 | 2 | 4 | 6 | 60
108AMXT-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
108AMXT-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
108LMXT | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 0.25 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 6 | 2 | 3 | 6 | 6 | 60
112AMXT-UPC1 | AMX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
112AMXT-UPC12 | AMX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
112LMXT | LMX | LMX | 230V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 0.5 | 1 | 2 | 3 | 5 | 6 | 8 | 10 | 5 | 5 | 10 | 6 | 60
120ASXT-UPC1 | ASX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 14 | 20 | 6 | 60
120ASXT-UPC12 | ASX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 14 | 20 | 6 | 60
125AMXT-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 150 | 1200 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
125AMXT-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
125LMXT | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 150 | 5000 | 1 | 2 | 5 | 8 | 10 | 12 | 15 | 20 | 5 | 10 | 20 | 9 | 90
140AMXT-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140AMXT-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140LMXT | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 15 | 150
140ASXT-UPC1 | ASX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 12 | 120
140ASXT-UPC12 | ASX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 2 | 5 | 8 | 10 | 15 | 20 | 30 | 5 | 16 | 30 | 12 | 120
160AMXT-UPC1 | AMX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 135 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160AMXT-UPC12 | AMX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160LMXT | LMX | LMX | 400V -6%/+10%, 50 Hz ±1% | 135 | 5000 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 15 | 150
160ASXT-UPC1 | ASX | UPC1 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 12 | 120
160ASXT-UPC12 | ASX | UPC12 | 400V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 1 | 5 | 10 | 15 | 30 | 35 | 40 | 48 | 5 | 16 | 48 | 12 | 120
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19

## 1Ph data (hidden)
Model | Type | Contr | Input | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current Single | MAX 2 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
115ASX-UPC1 | ASX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 |  | 2 | 15 | 4 | 40
115ASX-UPC12 | ASX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 |  | 2 | 15 | 4 | 40

## 1PhT data (hidden)
Model | Type | Contr | Input | V | Hz | Current 1 | Current 2 | Current 3 | Current 4 | Current 5 | Current 6 | Current Check | Current 8 | Current Single | MAX 2 Phase | MAX 1 Phase | Amps to Volts | Range 3 | Range 1
115ASXT-UPC1 | ASX | UPC1 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 |  | 2 | 15 | 4 | 40
115ASXT-UPC12 | ASX | UPC12 | 230V -6%/+10%, 50 Hz ±1% | 132 | 1200 | 0.5 | 1 | 2 | 4 | 6 | 8 | 10 | 15 |  | 2 | 15 | 4 | 40

## Other data do not use (hidden)
I. A. Norman
K. Gawlik
D. Holyoake
230V -6%/+10%, 50 Hz ±1%
400V -6%/+10%, 50 Hz ±1%
Results are recorded “As Received” values.
Results are recorded following repair and adjustments.
Results are recorded following adjustments.
ASX
AMX
ACX
ADX
AFX
LMX
UPC1
UPC12
UPC31
UPC32
UPC3
PRI
SEC
A

## Limits (hidden)
UPC1/3 |  |  | UPC12/32 | Amps to Volts
Frequency | Range | 15-1200 | 20-5000
| Accuracy | 0.01& of full sclae
400
1200
Voltage | Range | 0-150/341 | 0-150/375
| Accuracy | 0.5% of full scale
|  | 0.75 | 1.875
Voltmeter | Range | 0-354Vl-n, 708Vl-l
| Accuracy | 0.2% F.S | 0.25% of rdg
50-500 Hz + 0.1% FS
Ammeter
0

## All types (hidden)
Model | V | Hz | 2pV | contr | type
305AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
305AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
305LMX | 135 | 5000 | 270 | LMX | LMX
308AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
308AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
308LMX | 135 | 5000 | 270 | LMX | LMX
312AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
312AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
312LMX | 135 | 5000 | 270 | LMX | LMX
315ASX-UPC3 | 132 | 1200 | 264 | UPC3 | ASX
315ASX-UPC32 | 132 | 1200 | 264 | UPC32 | ASX
320AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
320AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
320LMX | 135 | 5000 | 270 | LMX | LMX
320ASX-UPC3 | 150 | 1200 | 300 | UPC3 | ASX
320ASX-UPC32 | 150 | 1200 | 300 | UPC32 | ASX
345AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
345AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
345LMX | 135 | 5000 | 270 | LMX | LMX
345ASX-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
345ASX-UPC32 | 135 | 5000 | 270 | UPC32 | ASX
360AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
360AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
360LMX | 135 | 5000 | 270 | LMX | LMX
360ASX-UPC3 | 132 | 1200 | 264 | UPC3 | ASX
360ASX-UPC32 | 132 | 1200 | 264 | UPC32 | ASX
390AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
390AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
390LMX | 135 | 5000 | 270 | LMX | LMX
390ASX-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
390ASX-UPC32 | 135 | 1200 | 270 | UPC32 | ASX
3120AMX-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
3120AMX-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
3120LMX | 135 | 5000 | 270 | LMX | LMX
3120ASX-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
3120ASX-UPC32 | 135 | 1200 | 270 | UPC32 | ASX
305AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
305AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
305LMXT | 135 | 5000 | 270 | LMX | LMX
308AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
308AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
308LMXT | 135 | 5000 | 270 | LMX | LMX
312AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
312AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
312LMXT | 135 | 5000 | 270 | LMX | LMX
315ASXT-UPC3 | 132 | 1200 | 264 | UPC3 | ASX
315ASXT-UPC32 | 132 | 1200 | 264 | UPC32 | ASX
320AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
320AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
320LMXT | 135 | 5000 | 270 | LMX | LMX
320ASXT-UPC3 | 150 | 1200 | 300 | UPC3 | ASX
320ASXT-UPC32 | 150 | 1200 | 300 | UPC32 | ASX
345AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
345AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
345LMXT | 135 | 5000 | 270 | LMX | LMX
345ASXT-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
345ASXT-UPC32 | 135 | 5000 | 270 | UPC32 | ASX
360AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
360AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
360LMXT | 135 | 5000 | 270 | LMX | LMX
360ASXT-UPC3 | 132 | 1200 | 264 | UPC3 | ASX
360ASXT-UPC32 | 132 | 1200 | 264 | UPC32 | ASX
390AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
390AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
390LMXT | 135 | 5000 | 270 | LMX | LMX
390ASXT-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
390ASXT-UPC32 | 135 | 1200 | 270 | UPC32 | ASX
3120AMXT-UPC3 | 135 | 1200 | 270 | UPC3 | AMX
3120AMXT-UPC32 | 135 | 5000 | 270 | UPC32 | AMX
3120LMXT | 135 | 5000 | 270 | LMX | LMX
3120ASXT-UPC3 | 135 | 1200 | 270 | UPC3 | ASX
3120ASXT-UPC32 | 135 | 1200 | 270 | UPC32 | ASX
360AFX | 300 | 1200 | 600 | AFX | AFX
390AFX | 300 | 5000 | 600 | AFX | AFX
3120AFX | 300 | 5000 | 600 | AFX | AFX
3150AFX | 300 | 1200 | 600 | AFX | AFX
3180AFX | 300 | 5000 | 600 | AFX | AFX
3240AFX | 300 | 5000 | 600 | AFX | AFX
3300AFX | 300 | 1200 | 600 | AFX | AFX
3450AFX | 300 | 5000 | 600 | AFX | AFX
3600AFX | 300 | 5000 | 600 | AFX | AFX
3750AFX | 300 | 1200 | 600 | AFX | AFX
3900AFX | 300 | 1200 | 600 | AFX | AFX
105AMX-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
105AMX-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
105LMX | 135 | 5000 | 270 | LMX | LMX
108AMX-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
108AMX-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
108LMX | 135 | 5000 | 270 | LMX | LMX
110ACX-UPC1 | 150 | 1200 | 300 | UPC1 | ACX
110ADX-UPC1 | 150 | 600 | 300 | UPC1 | ADX
112AMX-UPC1 | 150 | 1200 | 300 | UPC1 | AMX
112AMX-UPC12 | 150 | 5000 | 300 | UPC12 | AMX
112LMX | 150 | 5000 | 300 | LMX | LMX
115ACX-UPC1 | 150 | 1200 | 300 | UPC1 | ACX
115ADX-UPC1 | 150 | 600 | 300 | UPC1 | ADX
118ACX-UPC1 | 150 | 1200 | 300 | UPC1 | ACX
120ASX-UPC1 | 150 | 1200 | 300 | UPC1 | ASX
120ASX-UPC12 | 150 | 1200 | 300 | UPC12 | ASX
125AMX-UPC1 | 150 | 1200 | 300 | UPC1 | AMX
125AMX-UPC12 | 150 | 5000 | 300 | UPC12 | AMX
125LMX | 150 | 5000 | 300 | LMX | LMX
140AMX-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
140AMX-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
140LMX | 135 | 5000 | 270 | LMX | LMX
140ASX-UPC1 | 135 | 1200 | 270 | UPC1 | ASX
140ASX-UPC12 | 135 | 1200 | 270 | UPC12 | ASX
160AMX-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
160AMX-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
160LMX | 135 | 5000 | 270 | LMX | LMX
160ASX-UPC1 | 132 | 1200 | 264 | UPC1 | ASX
160ASX-UPC12 | 132 | 1200 | 264 | UPC12 | ASX
105AMXT-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
105AMXT-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
105LMXT | 135 | 5000 | 270 | LMX | LMX
108AMXT-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
108AMXT-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
108LMXT | 135 | 5000 | 270 | LMX | LMX
112AMXT-UPC1 | 150 | 1200 | 300 | UPC1 | AMX
112AMXT-UPC12 | 150 | 5000 | 300 | UPC12 | AMX
112LMXT | 150 | 5000 | 300 | LMX | LMX
120ASXT-UPC1 | 150 | 1200 | 300 | UPC1 | ASX
120ASXT-UPC12 | 150 | 1200 | 300 | UPC12 | ASX
125AMXT-UPC1 | 150 | 1200 | 300 | UPC1 | AMX
125AMXT-UPC12 | 150 | 5000 | 300 | UPC12 | AMX
125LMXT | 150 | 5000 | 300 | LMX | LMX
140AMXT-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
140AMXT-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
140LMXT | 135 | 5000 | 270 | LMX | LMX
140ASXT-UPC1 | 135 | 1200 | 270 | UPC1 | ASX
140ASXT-UPC12 | 135 | 1200 | 270 | UPC12 | ASX
160AMXT-UPC1 | 135 | 1200 | 270 | UPC1 | AMX
160AMXT-UPC12 | 135 | 5000 | 270 | UPC12 | AMX
160LMXT | 135 | 5000 | 270 | LMX | LMX
160ASXT-UPC1 | 132 | 1200 | 264 | UPC1 | ASX
160ASXT-UPC12 | 132 | 1200 | 264 | UPC12 | ASX
115ASX-UPC1 | 132 | 1200 | 264 | UPC1 | ASX
115ASX-UPC12 | 132 | 1200 | 264 | UPC12 | ASX
115ASXT-UPC1 | 132 | 1200 | 264 | UPC1 | ASX
115ASXT-UPC12 | 132 | 1200 | 264 | UPC12 | ASX
UPC1 | UPC3
UPC3 | UPC3
UPC12 | UPC32
UPC32 | UPC32
LMX | LMX
AFX | AFX